
 

NOVOBIOCIN DIFFERENTIATION DISKS

	Cat. no. Z7291
	Novobiocin Differentiation Disk, 5ug
	50 disks/cartridge


INTENDED USE

HardyDisk™ Novobiocin Differentiation Disks are useful for presumptively distinguishing Staphylococcus saprophyticus from other coagulase-negative staphylococci in clinical specimens.

SUMMARY

Coagulase-negative staphylococci (CoNS) are among those organisms that have traditionally been considered skin contaminants, and their recovery from cultures doesn't always indicate presence of disease. Therefore, little attention had been paid to the pathogenic potential of this group of bacteria until recently. By the mid-1970's, microbiologists were becoming aware that CoNS could indeed be pathogenic, especially in compromised hosts.(10) Today, S. saprophyticus has proven to be an important uropathogen.(9) It is second only to E. coli as the most common cause of cystitis and acute urinary tract infection (UTI) in healthy, young adult women.(5) S. saprophyticus tends to adhere to uroepithelial cells more often and more successfully than other staphylococcal species, this is believed to partially explain the organism's frequent role in urinary tract infections.(5,8)
The HardyDisk™ Novobiocin Differentiation Disks are useful in presumptively distinguishing S. saprophyticus from other CoNS. Other human staphylococcal species that are novobiocin-resistant (S. cohnii, S. xylosus, S. pulvereri) are rarely isolated from patients.(2-5) The novobiocin susceptibility test can be done using a plate method (18-24 hour test) or a tube method (5 hour test). A study conducted by Harrington and Gaydos in 1984, concluded that the novobiocin tube test is an acceptable method when performed using Tryptic Soy Broth (TSB), 3ml, and has the advantage of taking only 5 hours.(6) While the plate method is the more common approach of the two, both procedures are outlined in the section entitled "Procedure".

FORMULA

Each HardyDisk™ Novobiocin Differentiation Disk is prepared by impregnating 5ug of novobiocin onto high quality 6mm diameter filter paper disks.

STORAGE AND SHELF LIFE

Storage: Upon receipt store at -20 to +8 degrees C. away from direct light. The disks should not be used if there are any signs of deterioration, discoloration, or if the expiration date has passed. Protect from light, excessive heat, and moisture.

The expiration date applies to the product in its intact packaging when stored as directed.

PRECAUTIONS

This product is for in vitro diagnostic use only and is to be used only by adequately trained and qualified laboratory personnel. Observe approved biohazard precautions and aseptic techniques. All laboratory specimens should be considered infectious and handled according to "standard precautions". The "Guideline for Isolation Precautions" is available from the Centers of Disease Control and Prevention at www.cdc.gov/ncidod/dhqp/gl_isolation.html.

For additional information regarding specific precautions for the prevention of the transmission of all infectious agents from laboratory instruments and materials, and for recommendations for the management of exposure to infectious disease, refer to CLSI document M29.

Sterilize all biohazard waste before disposal.

Refer to the keyword "Precautions", in the Hardy Diagnostics software program HUGO™, for more information regarding general precautions when using culture media.

Refer to the keyword "MSDS", in the Hardy Diagnostics software program HUGO™, for more information on handling potentially hazardous material.

PROCEDURE

Specimen Collection: This product is not intended for primary isolation of patient specimens. It should be used only with cultures of isolated organisms. This product is used in conjunction with other biochemical tests to identify cultures of isolated organism.

Plate Method:

1. Allow disks to equilibrate to room temperature.

2. Using a pure 18-24 hour culture, prepare a suspension, equivalent to a McFarland 0.5 opacity standard, in Tryptic Soy Broth (TSB, Cat. no. R30), Sterile Water (Cat. no. U85), or Brain Heart Infusion (BHI) Broth (Cat. no. R15).

3. Inoculate Mueller Hinton Agar (Cat. no. G45), Blood Agar, 5% (Cat. no. A10), or Tryptic Soy Agar (TSA, Cat. no. G60) plate with a sterile swab to obtain confluent growth.

4. Aseptically apply one novobiocin disk onto the inoculated agar surface and lightly press down to ensure full contact with the medium.

5. Incubate aerobically for 18-24 hours at 35-37 degrees C.

6. Measure (in millimeters) the diameter of the zone of inhibition around the novobiocin disk, and record as susceptible or resistant. See section below, entitled "Interpretation of Results."

Tube Method:

1. Using a pure 18-24 hour culture, lightly inoculate two 3ml TSB tubes (Cat. no. R41) with the organism to be tested. The inoculum should be light, with no visible turbidity.

2. Immediately after inoculation, add one novobiocin disk to one of the tubes and shake for 5-10 seconds. The tube without the novobiocin disk will be the control tube.

3. Incubate both tubes aerobically at 35-37 degrees C. for up to 5 hours, or until the control tube (without novobiocin disk) shows turbidity comparable to that of a McFarland 0.5 opacity standard.

4. Check for turbidity in the tube containing the novobiocin disk as compared with the control tube, and record as susceptible or resistant. See section below, entitled "Interpretation of Results".

INTERPRETATION OF RESULTS

Plate Method:

Sensitive - A zone of inhibition greater than 16mm.(2)
Resistant - A zone of inhibition less than or equal to 16mm.(2)
Tube Method:

Sensitive - Turbidity is less than the control tube.
Resistant - Turbidity is more than or equal to the control tube.

LIMITATIONS

It is recommended that biochemical and/or serological tests be performed on colonies from pure culture for complete identification.

The novobiocin disk is not helpful and can give misleading results if it is performed on isolates other that those from urinary specimens.

Occasional human isolates that are not S. saprophyticus, S. cohnii subsp., or S. xylosis may also be resistant to novobiocin.(5)
Refer to the keyword "Limitations", in the Hardy Diagnostics software program HUGO™, for more information regarding general limitations on culture media.

MATERIALS REQUIRED BUT NOT PROVIDED

Standard microbiological supplies and equipment such as loops, other culture media, swabs, applicator sticks, incinerators, and incubators, etc., as well as serological and biochemical reagents, are not provided.

QUALITY CONTROL

The following organisms are routinely used for testing at Hardy Diagnostics:
	Test Organisms
	Inoculation Method*
	Incubation
	Results
(Plate/Tube Method)

	
	
	Time
	Temperature
	Atmosphere
	

	Staphylococcus saprophyticus
ATCC® 15305
	*
	*
	35-37°C
	Aerobic 
	Resistant -
Plate: Zone is less than or equal to 16mm
Tube: Turbidity is more than or equal to control tube

	Staphylococcus epidermidis
ATCC® 12228
	*
	*
	35-37°C
	Aerobic 
	Sensitive -
Plate: Zone is greater than 16mm
Tube: Turbidity is less than control tube


User Quality Control

Check for signs of contamination and deterioration. It is recommended that each new lot of disks be tested with known positive and negative controls and retested each week of use thereafter. Refer to the following keywords, in the Hardy Diagnostics software program HUGO™, for more information on QC: "Introduction to QC", "QC of Finished Product", and "The CLSI (NCCLS) Standard and Recommendations for User QC of Media". Also see listed references for more information.(3-6)
* Refer to section above, entitled "Procedures" for a description of inoculation procedures, incubation time and temperature.

PHYSICAL APPEARANCE

HardyDisk™ Novobiocin Differentiation Disks are 6mm (in diameter) filter paper disks with the letters NB5 printed on both sides and should appear white in color.


 
Novobiocin-resistant
(zone of inhibition < 16mm).
Staphylococcus saprophyticus (ATCC® 15305) growing around a HardyDisk™ Novobiocin Differentiation Disk (Cat. no. Z7291). Incubated aerobically on TSA (Cat. no. G60) for 24 hours at 35 deg. C.

 
Novobiocin-sensitive
(zone of inhibition > 16mm).
Staphylococcus epidermidis (ATCC® 12228) inhibition zone around a HardyDisk™ Novobiocin Differentiation Disk (Cat. no. Z7291). Incubated aerobically on TSA (Cat. no. G60) for 24 hours at 35 deg. C.

REFERENCES

1. August, M.J., et al. 1990. Cumitech 3A; Quality Control and Quality Assurance Practices in Clinical Microbiology, Coordinating ed., A.S. Weissfeld. American Society for Microbiology, Washington, D.C.

2. Murray, P.R., et al. 2003. Manual of Clinical Microbiology, 8th ed. American Society for Microbiology, Washington, D.C.

3. Baron, Ellen Jo, Ph.D., et al. 1994. Bailey and Scott's Diagnostic Microbiology, 9th ed. C.V. Mosby Company, St. Louis, MO.

4. Isenberg, H.D. Clinical Microbiology Procedures Handbook, Vol. I & II. American Society for Microbiology, Washington, D.C.

5. Koneman, E.W., et al. 1997. Color Atlas and Textbook of Diagnostic Microbiology, 5th ed. J.B. Lippincott Company, Philadelphia, PA.

6. Harrington, Brian J., and J.M. Gaydos. 1984. Five-Hour Novobiocin Test for Differentiation of Coagulase-Negative Staphylococci. J. Clin. Microbiol.; 19:279-280.

7. Gregson, Dan B., D.E. Low, M. Skulnick, and A.E. Simor. 1988. Problems with Rapid Agglutination Methods for Identification of Staphylococcus aureus When Staphylococcus saprophyticus Is Being Tested. J. Clin. Microbiol.; 26:1398-1399.

8. Sneath, P.H.A., N.S. Mair, M.E. Sharpe, J.G. Holt. 1986. Section 12. Gram-Positive Cocci, p. 1026-1027. Bergy's Manual of Systematic Bacteriology, Vol. II. Williams & Wilkins, Baltimore, MD.

9. Rupp, M.E., D.E. Soper, and G.L. Archer. 1992. Colonization of the female Genital Tract with Staphylococcus saprophyticus. J. Clin. Microbiol.; 30:2975-2979.

10. Conville, P.S., Ann C. Albers. 1984. Staphylococcus saprophyticus: A Review of the Literature. Journal of Medical Technology.; 1:6 p. 513-519.

ATCC is a registered trademark of the American Type Culture Collection.

062006md

HARDY DIAGNOSTICS
1430 West McCoy Lane, Santa Maria, CA 93455, USA
Phone: (805) 346-2766 ext. 5658
Fax: (805) 346-2760
Website: www.HardyDiagnostics.com
Email: TechService@HardyDiagnostics.com
Ordering Information
Distribution Centers:
California · Washington · Utah · Arizona · Texas · Ohio · New York · Florida

The Hardy Diagnostics manufacturing facility and quality
management system is certified to ISO 13485.
Copyright© 1996 - 2013 by Hardy Diagnostics. All rights reserved.
